## Slide 1: OCN 479-001: Smart Coasts
Week 7
Power Budget, Sensor specs, Metadata vs datadata

## Slide 2: Power Budgeting
How do you figure out how long something will last in the field?
How much capacity you have/how much sensor(s) will draw

## Slide 3: Budget concept
General:
  Available = initial + additions – subtractions
Finance:
  Balance = savings + earnings – expenditures
Power:
  ?

## Slide 4: Example sensor:
SenseAir K-30, a low-cost indoor air quality CO2 sensor
https://www.instructables.com/Track-CO2-Carbon-Dioxide-Levels-With-Ardunio-and-C/

## Slide 5: Power Budget for Sensor
Step 1: look up component datasheets, e.g.,

## Slide 6: 3.7 V “nominal”
https://community.particle.io/t/can-argon-or-xenon-read-the-battery-state/45554/45?u=fragma
https://www.adafruit.com/product/258
Step 2: look up how much power you have, multiply A-h by V to get W-h

## Slide 7: Simple Estimate
We have 1200 mA*hr * 3.7 V = 4.44 W*hr
We will use > 40 mA * 5 V (for K-30; ignore microcontroller for now) = 200 mW = 0.2 W

## Slide 8: Simple Estimate
We have 1200 mA*hr * 3.7 V = 4.44 W*hr
We will use > 40 mA * 5 V (for K-30; ignore microcontroller for now) = 200 mW = 0.2 W
So 4.44 W*hr/0.2 W = 22 hours

## Slide 9: Taking it further: measure current in all states (e.g., sampling, sleeping, publishing)
https://www.sciencebuddies.org/science-fair-projects/references/how-to-use-a-multimeter

## Slide 10: Sensor Specifications

## Slide 11: Important Metrics
Match a letter from the left to the following:
  Accurate but imprecise
  Precise but inaccurate
  Imprecise and inaccurate
  Precise and accurate

## Slide 12: Important Metrics
Match a letter from the left to the following:
  Accurate but imprecise
  Precise but inaccurate
  Imprecise and inaccurate
  Precise and accurate
C

## Slide 13: Important Metrics
Match a letter from the left to the following:
  Accurate but imprecise
  Precise but inaccurate
  Imprecise and inaccurate
  Precise and accurate
C
B

## Slide 14: Important Metrics
Match a letter from the left to the following:
  Accurate but imprecise
  Precise but inaccurate
  Imprecise and inaccurate
  Precise and accurate
C
B
A
D

## Slide 15: Important Metrics
Example: ”true” temperature is 100 ºC
Take 10 samples. Get average and stdev (report as average±stdev)
Accurate but imprecise: thermometer measures 100.1±36 ºC
Inaccurate but precise: thermometer measures 56±0.2 ºC
Inaccurate, imprecise: 56±36 ºF
Accurate, precise: 100.1±0.2 ºC

## Slide 16: Important Metrics
Example: ”true” temperature is 100 ºF
Take 10 samples. Get average and stdev (report as average±stdev)
Accurate but imprecise: thermometer measures 100.1±36 ºC
Inaccurate but precise: thermometer measures 56±0.2 ºC
Inaccurate, imprecise: 56±36 ºF
Accurate, precise: 100.1±0.2 ºC
Extremely important: “accuracy” and “precision” are relative terms. Values ascribed to them depends on application!

## Slide 17: Important Metrics
Accuracy, precision
https://manoa.hawaii.edu/exploringourfluidearth/physical/world-ocean/map-distortion/practices-science-precision-vs-accuracy

## Slide 18: Important Metrics
Accuracy, precision, stability
Continuing previous example:
  Actual temperature not changing, but measurements are:
    t = 0 min, T = 99.8±0.2 ºC
    t = 10 min, T = 101.4±0.1 ºC
    t = 20 min, T = 103.7±0.2 ºC
https://manoa.hawaii.edu/exploringourfluidearth/physical/world-ocean/map-distortion/practices-science-precision-vs-accuracy
t = 0 min
t = 10 min
t = 20 min

## Slide 19: Important Metrics
Accuracy, precision, stability, resolution
Does your thermometer tell you temperature to 1 degree? 0.1? 0.01?
Just because there are more decimal points (higher resolution) does not mean it is more accurate!
https://www.allsensors.com/engineering-resources/white-papers/accuracy-and-precision-for-mems-pressure-sensors

## Slide 20: Important Metrics
Accuracy, precision, stability, resolution
Does your thermometer tell you temperature to 1 degree? 0.1? 0.01?
Just because there are more decimal points (higher resolution) does not mean it is more accurate!
https://www.researchgate.net/figure/Quantization-resolution-in-this-example-is-4_fig2_352898102

## Slide 21: Important Metrics
Accuracy/precision/stability/resolution
Response time (imagine you move a thermometer from an icewater bucket to a pot of boiling water; how long will it take for the thermometer to change temps?)
https://blog.wika.us/products/temperature-products/temperature-sensors-thermowells-and-response-times/

## Slide 22: Response time defined
https://www.baranidesign.com/faq-articles/2019/5/6/difference-between-sensor-response-time-and-sensor-time-constant-tau

## Slide 23: Available technologies
Which value do sensor manufacturers use? (Whatever they want—zero consistency in this arena)
If your project requires accuracy, precision, stability, response time, etc., VERIFY for yourself
Some manufacturers very honest & thorough results, some report “best case scenario” or results of one specific, narrowly defined test

## Slide 24: Spec sheet example
What is accuracy @ current atmospheric CO2?
Just because it works over a range of temp, pressure, salinity, etc. doesn’t mean it performs the same!

## Slide 25: Spec sheet example

## Slide 26: Data and Metadata

## Slide 27: Data vs. Metadata
Data:
  A measurement (from a sensor, analyzer, a ruler, etc.)
  Observations, including numbers and words

## Slide 28: Data vs. Metadata
Data:
  A measurement (from a sensor, analyzer, a ruler, etc.)
  Observations, including numbers and words
Metadata
  Data about the data
  How were data collected?
  What was the instrument?
  When was it last calibrated?
  What are its calibration coefficients?
  Who made the measurement?

## Slide 29: Why?
Research is meant to be reproducible
  By you
  By anyone who reads your paper, lab notebook, etc.
You’ll remember very little of your past work after weeks/months/years

## Slide 31
METADATA
DATA
COLUMN HEADERS